Measurement Protocol PAX

# General Information:

|  |  |  |
| --- | --- | --- |
| Test date: 07/08/20 | Date prev. test: | Test number: 1 |
| Tested by: Sarah Schoultz | PAX number: PB-004 | Installed at ant.: |
| Comments: | | |
|  | | |
|  | | |
|  | | |
| Known problems: | | |
|  | | |
|  | | |
|  | | |

# Setup:

|  |  |  |  |
| --- | --- | --- | --- |
| Power supply: | OK 🗹 NOK 🗆 | Communication: | OK 🗹 NOK 🗆 |
| Supply current +6V: 1.125mA | | ssh obs@antcntl | |
| Supply current -6V: 0.117mA | | ssh ataant@paxtester | pw: q@n@t |
| Supply current +5V: 0.245mA | | telnet pax | “help” |
|  | | CTRL + ] | “close” |

# RF Test with VNA:

### VNA Setup:

|  |  |  |
| --- | --- | --- |
| Load configuration: Passband\_PAX.csa | | |
| Freq Start: 1.0 GHz | Freq Stop: 13.0 GHz | Power Level: -20 dBm |
| Averaging: enabled | Averaging count: 15 | N. Points: 801 |
| VNA : Agilent N5230C 10MHz - 20GHz | | |

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| VNA to PAX connection: | 1m ABC-CA18 Cable + 20dB Attenuator | | | |
| Input Power Level to PAX: (Measured) | @1GHz:  -40.2dBm | @4GHz:  -40.7dBm | @8GHz: -41.2dBm | @12GHz: -41.4dBm |

|  |  |
| --- | --- |
| PAX to VNA connection: | 2m Fibre Cable + Fibre Diode + AOX Amplifier + 1m ABC-CA18 Cable |

### Measured Passband with Attenuator set to 7dB each (Complete Link):

|  |
| --- |
| X |

A screenshot of a cell phone

Description automatically generated

|  |  |  |  |
| --- | --- | --- | --- |
| Flatness: | OK 🗹 NOK 🗆 | Unwanted Features: | Yes 🗆 No 🗹 |

|  |
| --- |
| Y |

A screenshot of a cell phone

Description automatically generated

|  |  |  |  |
| --- | --- | --- | --- |
| Flatness: | OK 🗹 NOK 🗆 | Unwanted Features: | Yes 🗆 No 🗹 |

### Power Sweep (Complete Link):

|  |  |  |
| --- | --- | --- |
| Load configuration: Power\_Sweep\_PAX.csa | | |
| Power Start: -40dBm | Power Stop: -20dBm | Frequency : 4.0GHz |
| PAM Attenuator: 0dB |  | N. Points: 801 |

|  |
| --- |
| X |

A screenshot of a cell phone

Description automatically generated

|  |  |  |  |
| --- | --- | --- | --- |
| Compression point: | OK 🗹 NOK 🗆 | Unwanted Features: | Yes 🗆 No 🗹 |

|  |
| --- |
| Y |

A screenshot of a cell phone

Description automatically generated

|  |  |  |  |
| --- | --- | --- | --- |
| Compression point: | OK 🗹 NOK 🗆 | Unwanted Features: | Yes 🗆 No 🗹 |

### Detector Calibration and Attenuator Sweep Pol X:

|  |  |  |
| --- | --- | --- |
| Load configuration: Detector\_Calibration\_PAX.csa | | |
| Freq : 4.0 GHz | Power Level: -20dBm | Sweep Time: -20sec |
| Sweep Mode: CW | Power Level at PAX Input (Measured): -40.6dBm | |

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| X | | | | |
| CW Input Power [dBm]: | Attenuator Value A [dB]: | Attenuator Value: B [dB]: | Detector Value: | Power Meter  Value [dBm]: |
| -40.6 | 0 | 0 | - | 14.3 |
| -40.6 | 0 | 3 | - | 14.0 |
| -40.6 | 0 | 6 | - | 11.9 |
| -40.6 | 0 | 9 | 0.8607 | 9.1 |
| -40.6 | 0 | 12 | 0.6636 | 6.1 |
| -40.6 | 0 | 15 | 0.4139 | 3.3 |
| -40.6 | 0 | 18 | 0.2331 | 0.1 |
| -40.6 | 0 | 21 | 0.1330 | -3.0 |
| -40.6 | 3 | 21 | 0.0673 | -6.3 |
| -40.6 | 6 | 21 | 0.0354 | -9.2 |
| -40.6 | 9 | 21 | 0.0187 | -12.2 |
| -40.6 | 12 | 21 | 0.0101 | -15.1 |
| -40.6 | 15 | 21 | 0.0062 | -18.0 |
| -40.6 | 18 | 21 | 0.0035 | -21.2 |
| -40.6 | 21 | 21 | 0.0024 | -24.3 |
| -40.6 | 24 | 21 | 0.0022 | -27.1 |
| -40.6 | 27 | 21 | 0.0016 | -30.1 |
| -40.6 | 30 | 21 | 0.0020 | -33.1 |
| -40.6 | 30 | 24 | 0.0012 | -35.6 |
| -40.6 | 30 | 27 | 0.0016 | -38.0 |
| -40.6 | 30 | 30 | 0.0015 | -40.1 |

|  |  |  |
| --- | --- | --- |
| Use Noise Source: Atlantic AS6333 | | |
| Freq : 1.0 - 12.0GHz | Power Level: -41.8dBm |  |
| DC Supply: 28V |  | |

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| X | | | | |
| Input Power [dBm]: | Attenuator Value A [dB]: | Attenuator Value: B [dB]: | Detector Value: | Power Meter  Value [dBm]: |
| -41.8 | 0 | 0 | - | 11.9 |
| -41.8 | 0 | 3 | - | 10.6 |
| -41.8 | 0 | 6 | - | 8.7 |
| -41.8 | 0 | 9 | - | 6.3 |
| -41.8 | 0 | 12 | 0.8598 | 3.6 |
| -41.8 | 0 | 15 | 0.6727 | 0.9 |
| -41.8 | 0 | 18 | 0.3864 | -2.2 |
| -41.8 | 0 | 21 | 0.2263 | -5.0 |
| -41.8 | 3 | 21 | 0.1273 | -8.0 |
| -41.8 | 6 | 21 | 0.0671 | -11.0 |
| -41.8 | 9 | 21 | 0.0359 | -13.9 |
| -41.8 | 12 | 21 | 0.0191 | -16.8 |
| -41.8 | 15 | 21 | 0.0112 | -19.5 |
| -41.8 | 18 | 21 | 0.0061 | -22.7 |
| -41.8 | 21 | 21 | 0.0043 | -25.7 |
| -41.8 | 24 | 21 | 0.0030 | -28.6 |
| -41.8 | 27 | 21 | 0.0024 | -31.4 |
| -41.8 | 30 | 21 | 0.0022 | -34.4 |
| -41.8 | 30 | 24 | 0.0018 | -36.8 |
| -41.8 | 30 | 27 | 0.0021 | -39.9 |
| -41.8 | 30 | 30 | 0.0019 | -40.7 |

### Detector Calibration and Attenuator Sweep Pol Y:

|  |  |  |
| --- | --- | --- |
| Load configuration: Detector\_Calibration\_PAX.csa | | |
| Freq : 4.0 GHz | Power Level: -20dBm | Sweep Time: -20sec |
| Sweep Mode: CW | Power Level at PAX Input (Measured): -40.6dBm | |

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| Y | | | | |
| CW Input Power [dBm]: | Attenuator Value A [dB]: | Attenuator Value: B [dB]: | Detector Value: | Power Meter  Value [dBm]: |
| -40.6 | 0 | 0 | - | 13.3 |
| -40.6 | 0 | 3 | - | 13.3 |
| -40.6 | 0 | 6 | - | 12.6 |
| -40.6 | 0 | 9 | - | 10.5 |
| -40.6 | 0 | 12 | - | 7.6 |
| -40.6 | 0 | 15 | 0.8577 | 4.8 |
| -40.6 | 0 | 18 | 0.5780 | 1.5 |
| -40.6 | 0 | 21 | 0.3341 | -1.6 |
| -40.6 | 3 | 21 | 0.1716 | -5.0 |
| -40.6 | 6 | 21 | 0.0921 | -8.0 |
| -40.6 | 9 | 21 | 0.0476 | -11.0 |
| -40.6 | 12 | 21 | 0.0251 | -14.0 |
| -40.6 | 15 | 21 | 0.0131 | -17.0 |
| -40.6 | 18 | 21 | 0.0065 | -20.2 |
| -40.6 | 21 | 21 | 0.0036 | -23.3 |
| -40.6 | 24 | 21 | 0.0024 | -26.2 |
| -40.6 | 27 | 21 | 0.0014 | -29.3 |
| -40.6 | 30 | 21 | 0.0011 | -32.55 |
| -40.6 | 30 | 24 | 0.0007 | -35.2 |
| -40.6 | 30 | 27 | 0.0008 | -37.7 |
| -40.6 | 30 | 30 | 0.0008 | -40.0 |

|  |  |  |
| --- | --- | --- |
| Use Noise Source: Atlantic AS6333 | | |
| Freq : 1.0 - 12.0GHz | Power Level: -41.8dBm |  |
| DC Supply: 28V |  | |

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| Y | | | | |
| Input Power [dBm]: | Attenuator Value A [dB]: | Attenuator Value: B [dB]: | Detector Value: | Power Meter  Value [dBm]: |
| -41.8 | 0 | 0 | - | 11.8 |
| -41.8 | 0 | 3 | - | 10.8 |
| -41.8 | 0 | 6 | - | 9.3 |
| -41.8 | 0 | 9 | - | 7.3 |
| -41.8 | 0 | 12 | - | 4.7 |
| -41.8 | 0 | 15 | 0.8530 | 2.2 |
| -41.8 | 0 | 18 | 0.4809 | -0.9 |
| -41.8 | 0 | 21 | 0.2724 | -3.9 |
| -41.8 | 3 | 21 | 0.1454 | -6.8 |
| -41.8 | 6 | 21 | 0.0742 | -9.9 |
| -41.8 | 9 | 21 | 0.0390 | -12.8 |
| -41.8 | 12 | 21 | 0.0198 | -15.9 |
| -41.8 | 15 | 21 | 0.0113 | -18.5 |
| -41.8 | 18 | 21 | 0.0058 | -21.8 |
| -41.8 | 21 | 21 | 0.0035 | -24.9 |
| -41.8 | 24 | 21 | 0.0023 | -27.9 |
| -41.8 | 27 | 21 | 0.0016 | -30.7 |
| -41.8 | 30 | 21 | 0.0015 | -34.0 |
| -41.8 | 30 | 24 | 0.0014 | -36.6 |
| -41.8 | 30 | 27 | 0.0016 | -38.7 |
| -41.8 | 30 | 30 | 0.0010 | -40.8 |

# Inspection:

### Visual:

|  |  |  |
| --- | --- | --- |
| Fibre Connectors Clean | Comments: | OK 🗹 NOK 🗆 |
| Fibre Connector Mechanical | Comments: | OK **🗹** NOK **🗆** |
| RF Cable | Comments: | OK **🗹** NOK **🗆** |
| RF Connectors Clean | Comments: | OK **🗹** NOK **🗆** |

### Functions:

|  |  |  |
| --- | --- | --- |
| Read Temperature Value | Comments: 34.8 | OK 🗹 NOK 🗆 |
| LNA Settings | Comments: | OK**🗹** NOK **🗆** |
| Temperature Stabilisation | Comments: | OK 🗆 NOK **🗆** |
|  |  | OK 🗆 NOK **🗆** |