# Measurement Protocol PAX

# General Information:

| Test date: Nov. 13, 2025 | Date prev. test: Feb 14 2020 | Test number: 1 |
| --- | --- | --- |
| Tested by: Alex Pollak/Karen | PAX number: PB-005 | Installed at ant.: 2B |
| Comments: Prototyping, replace new detector assembly in Y-pol |  |  |
| Known problems: |  |  |

# Setup:

| Power supply: | OK NOK | Communication: | OK NOK |
| --- | --- | --- | --- |
| Supply current +6V: 1.04A |  | ssh obs@antcntl |  |
| Supply current -6V: 0.075A |  | ssh ataant@paxtester | pw: q@n@t |
| Supply current +5V: 0.044A |  | telnet pax | “help” |
|  |  | CTRL + ] | “close” |

# RF Test with VNA:

### VNA Setup:

| Load configuration: Passband_PAX.sta |  |  |
| --- | --- | --- |
| Freq Start: 1.0 GHz | Freq Stop: 13.0 GHz | Power Level: -20 dBm |
| Averaging: enabled | Averaging count: 15 | N. Points: 801 (under Sweep) |
| VNA : Agilent N5230C 10MHz - 20GHz |  |  |

| VNA to PAX connection: | 1m ABC-CA18 Cable + 20dB Attenuator |  |  |  |
| --- | --- | --- | --- | --- |
| Input Power Level to PAX: (Measured) | @1GHz: -40.2dBm | @4GHz: -40.7dBm | @8GHz: -41.2dBm | @12GHz: -41.4dBm |

| PAX to VNA connection: | 2m Fibre Cable + Lab Buddy + AOX Amplifier + 1m ABC-CA18 Cable |
| --- | --- |

### Measured Passband with Attenuator set to 7dB each (Complete Link):

X

| Flatness: | OK NOK | Unwanted Features: | Yes No |
| --- | --- | --- | --- |

Y

| Flatness: | OK NOK | Unwanted Features: | Yes No |
| --- | --- | --- | --- |

Power Sweep (Complete Link):

| Load configuration: Power_Sweep_PAX.sta |  |  |
| --- | --- | --- |
| Power Start: -40dBm | Power Stop: -20dBm | Frequency : 4.0GHz |
| PAM Attenuator: 0dB |  | N. Points: 801 |

X

| Compression point: | OK NOK | Unwanted Features: | Yes No |
| --- | --- | --- | --- |

Y

| Compression point: | OK NOK | Unwanted Features: | Yes No |
| --- | --- | --- | --- |

### Detector Calibration and Attenuator Sweep Pol X:

Use Signal generator instead of VNA now

SMB-100B signal generator

Remove 20dB attenuator

Use 2m ABC-CA18 cable

Connect to diode power sensor

| Load configuration: Detector_Calibration_PAX.csa |  |  |
| --- | --- | --- |
| Freq : 4.0 GHz | Power Level: -38.2dBm |  |
| Mode: CW (Continuous Wave) | Power Level at PAX Input (Measured): -40.1dBm |  |

| X |  |  |  |  |
| --- | --- | --- | --- | --- |
| CW Input Power [dBm]: | Attenuator Value A [dB]: | Attenuator Value: B [dB]: | Detector Value: | Power Meter Value [dBm]: |
| -40.1 | 0 | 0 | 0.8377 | 12.4 |
| -40.1 | 0 | 3 | 0.8381 | 12.3 |
| -40.1 | 0 | 6 | 0.8399 | 11.6 |
| -40.1 | 0 | 9 | 0.8364 | 9.1 |
| -40.1 | 0 | 12 | 0.8383 | 6.7 |
| -40.1 | 0 | 15 | 0.8396 | 4.4 |
| -40.1 | 0 | 18 | 0.6669 | 0.3 |
| -40.1 | 0 | 21 | 0.3908 | -2.1 |
| -40.1 | 3 | 21 | 0.1905 | -5.4 |
| -40.1 | 6 | 21 | 0.0932 | -8.3 |
| -40.1 | 9 | 21 | 0.0531 | -11.4 |
| -40.1 | 12 | 21 | 0.0268 | -14.5 |
| -40.1 | 15 | 21 | 0.0125 | -17.5 |
| -40.1 | 18 | 21 | 0.0074 | -20.5 |
| -40.1 | 21 | 21 | 0.0042 | -23.4 |
| -40.1 | 24 | 21 | 0.0026 | -26.1 |
| -40.1 | 27 | 21 | 0.0017 | -29.2 |
| -40.1 | 30 | 21 | 0.0012 | -32.4 |
| -40.1 | 30 | 24 | 0.0011 | -35.1 |
| -40.1 | 30 | 27 | 0.0010 | -37.5 |
| -40.1 | 30 | 30 | 0.0009 | -40.0 |

| Use Noise Source: Atlantic AS6333 |  |  |
| --- | --- | --- |
| Freq : 1.0 - 12.0GHz | Power Level: -41.8dBm |  |
| DC Supply: 28V |  |  |

| X |  |  |  |  |
| --- | --- | --- | --- | --- |
| Input Power [dBm]: | Attenuator Value A [dB]: | Attenuator Value: B [dB]: | Detector Value: | Power Meter Value [dBm]: |
| -41.8 | 0 | 0 | 0.8325 | 9.7 |
| -41.8 | 0 | 3 | 0.8349 | 9.1 |
| -41.8 | 0 | 6 | 0.8364 | 7.6 |
| -41.8 | 0 | 9 | 0.8379 | 5.9 |
| -41.8 | 0 | 12 | 0.8374 | 3.3 |
| -41.8 | 0 | 15 | 0.8165 | 1.2 |
| -41.8 | 0 | 18 | 0.4429 | -1.9 |
| -41.8 | 0 | 21 | 0.2533 | -4.5 |
| -41.8 | 3 | 21 | 0.1079 | -8.5 |
| -41.8 | 6 | 21 | 0.0574 | -11.1 |
| -41.8 | 9 | 21 | 0.0278 | -13.9 |
| -41.8 | 12 | 21 | 0.0160 | -16.5 |
| -41.8 | 15 | 21 | 0.0090 | -19.3 |
| -41.8 | 18 | 21 | 0.0042 | -22.6 |
| -41.8 | 21 | 21 | 0.0024 | -26.1 |
| -41.8 | 24 | 21 | 0.0017 | -29.2 |
| -41.8 | 27 | 21 | 0.0013 | -31.7 |
| -41.8 | 30 | 21 | 0.0011 | -35.8 |
| -41.8 | 30 | 24 | 0.0009 | -37.2 |
| -41.8 | 30 | 27 | 0.0009 | -39.3 |
| -41.8 | 30 | 30 | 0.0009 | -41.3 |

### Detector Calibration and Attenuator Sweep Pol Y:

| Load configuration: Detector_Calibration_PAX.csa |  |  |
| --- | --- | --- |
| Freq : 4.0 GHz | Power Level: -20dBm | Sweep Time: -20sec |
| Sweep Mode: CW | Power Level at PAX Input (Measured): -40.1dBm |  |

| Y |  |  |  |  |
| --- | --- | --- | --- | --- |
| CW Input Power [dBm]: | Attenuator Value A [dB]: | Attenuator Value: B [dB]: | Detector Value: | Power Meter Value [dBm]: |
| -40.1 | 0 | 0 | 0.1899 | 10.8 |
| -40.1 | 0 | 3 | 0.1862 | 10.5 |
| -40.1 | 0 | 6 | 0.1694 | 9.5 |
| -40.1 | 0 | 9 | 0.1411 | 7.6 |
| -40.1 | 0 | 12 | 0.1058 | 4.8 |
| -40.1 | 0 | 15 | 0.0802 | 2.1 |
| -40.1 | 0 | 18 | 0.0561 | -1.5 |
| -40.1 | 0 | 21 | 0.0432 | -4.5 |
| -40.1 | 3 | 21 | 0.0226 | -7.8 |
| -40.1 | 6 | 21 | 0.0163 | -12.9 |
| -40.1 | 9 | 21 | 0.0119 | -15.1 |
| -40.1 | 12 | 21 | 0.0097 | -18.3 |
| -40.1 | 15 | 21 | 0.0072 | -21.7 |
| -40.1 | 18 | 21 | 0.0054 | -24.6 |
| -40.1 | 21 | 21 | 0.0044 | -27.6 |
| -40.1 | 24 | 21 | 0.0038 | -30.6 |
| -40.1 | 27 | 21 | 0.0036 | -32.8 |
| -40.1 | 30 | 21 | 0.0033 | -35.4 |
| -40.1 | 30 | 24 | 0.0031 | -37.6 |
| -40.1 | 30 | 27 | 0.0030 | -39.5 |
| -40.1 | 30 | 30 | 0.0029 | -42.0 |

| Use Noise Source: Atlantic AS6333 |  |  |
| --- | --- | --- |
| Freq : 1.0 - 12.0GHz | Power Level: -41.8dBm |  |
| DC Supply: 28V |  |  |

| Y |  |  |  |  |
| --- | --- | --- | --- | --- |
| Input Power [dBm]: | Attenuator Value A [dB]: | Attenuator Value: B [dB]: | Detector Value: | Power Meter Value [dBm]: |
| -41.8 | 0 | 0 | 0.1735 | 8.4 |
| -41.8 | 0 | 3 | 0.1437 | 7.3 |
| -41.8 | 0 | 6 | 0.1205 | 4.6 |
| -41.8 | 0 | 9 | 0.0947 | 2.6 |
| -41.8 | 0 | 12 | 0.0631 | -1.3 |
| -41.8 | 0 | 15 | 0.0497 | -3.6 |
| -41.8 | 0 | 18 | 0.0338 | -5.5 |
| -41.8 | 0 | 21 | 0.0259 | -8.5 |
| -41.8 | 3 | 21 | 0.0187 | -11.2 |
| -41.8 | 6 | 21 | 0.0138 | -14.2 |
| -41.8 | 9 | 21 | 0.0098 | -16.9 |
| -41.8 | 12 | 21 | 0.0073 | -20.0 |
| -41.8 | 15 | 21 | 0.0065 | -22.6 |
| -41.8 | 18 | 21 | 0.0045 | -25.8 |
| -41.8 | 21 | 21 | 0.0039 | -28.9 |
| -41.8 | 24 | 21 | 0.0034 | -31.8 |
| -41.8 | 27 | 21 | 0.0031 | -34.6 |
| -41.8 | 30 | 21 | 0.0029 | -37.8 |
| -41.8 | 30 | 24 | 0.0029 | -40.2 |
| -41.8 | 30 | 27 | 0.0028 | -42.3 |
| -41.8 | 30 | 30 | 0.0029 | -44.4 |

# Inspection:

### Visual:

| Fibre Connectors Clean | Comments: | OK NOK |
| --- | --- | --- |
| Fibre Connector Mechanical | Comments: | OK NOK |
| RF Cable | Comments: | OK NOK |
| RF Connectors Clean | Comments: | OK NOK |

### Functions:

| Read Temperature Value | Comments: 31.7 | OK NOK |
| --- | --- | --- |
| LNA Settings | Comments: | OK NOK |
| Temperature Stabilisation | Comments: | OK NOK |
|  |  | OK NOK |