# Measurement Protocol PAX

# General Information:

| Test date: 06/11/26 | Date prev. test: 13/08/20 | Test number: 2 |
| --- | --- | --- |
| Tested by: Jay Frothingham | PAX number: PB-043 | Installed at ant.: |
| Comments: |  |  |
| Replaced X-pam from 009, remeasurement. |  |  |
| Known problems: |  |  |

# Setup:

| Power supply: | OK NOK | Communication: | OK NOK |
| --- | --- | --- | --- |
| Supply current +6V: 1.084A |  | ssh obs@antcntl |  |
| Supply current -6V: 0.109A |  | ssh ataant@paxtester | pw: q@n@t |
| Supply current +5V: 0.045A |  | telnet pax | “help” |
|  |  | CTRL + ] | “close” |

# RF Test with VNA:

### VNA Setup:

| Load configuration: Passband_PAX.csa |  |  |
| --- | --- | --- |
| Freq Start: 1.0 GHz | Freq Stop: 13.0 GHz | Power Level: -20 dBm |
| Averaging: enabled | Averaging count: 15 | N. Points: 801 |
| VNA : Agilent N5230C 10MHz - 20GHz |  |  |

| VNA to PAX connection: | 1m ABC-CA18 Cable + 20dB Attenuator |  |  |  |
| --- | --- | --- | --- | --- |
| Input Power Level to PAX: (Measured) | @1GHz: -40.2dBm | @4GHz: -40.7dBm | @8GHz: -41.2dBm | @12GHz: -41.4dBm |

| PAX to VNA connection: | 2m Fibre Cable + Fibre Diode + AOX Amplifier + 1m ABC-CA18 Cable |
| --- | --- |

### Measured Passband with Attenuator set to 7dB each (Complete Link):

X

| Flatness: | OK NOK | Unwanted Features: | Yes No |
| --- | --- | --- | --- |

Y

| Flatness: | OK NOK | Unwanted Features: | Yes No |
| --- | --- | --- | --- |

### Power Sweep (Complete Link):

| Load configuration: Power_Sweep_PAX.csa |  |  |
| --- | --- | --- |
| Power Start: -40dBm | Power Stop: -20dBm | Frequency : 4.0GHz |
| PAM Attenuator: 0dB |  | N. Points: 801 |

X

| Compression point: | OK NOK | Unwanted Features: | Yes No |
| --- | --- | --- | --- |

Y

| Compression point: | OK NOK | Unwanted Features: | Yes No |
| --- | --- | --- | --- |

Detector Calibration and Attenuator Sweep Pol X:

| Load configuration: Detector_Calibration_PAX.csa |  |  |
| --- | --- | --- |
| Freq : 4.0 GHz | Power Level: -20dBm | Sweep Time: -20sec |
| Sweep Mode: CW | Power Level at PAX Input (Measured): -40.6dBm |  |

| X |  |  |  |  |
| --- | --- | --- | --- | --- |
| CW Input Power [dBm]: | Attenuator Value A [dB]: | Attenuator Value: B [dB]: | Detector Value: | Power Meter Value [dBm]: |
| -40.6 | 0 | 0 | - | 12.9 |
| -40.6 | 0 | 3 | - | 12.9 |
| -40.6 | 0 | 6 | - | 12.3 |
| -40.6 | 0 | 9 | - | 10.3 |
| -40.6 | 0 | 12 | - | 7.4 |
| -40.6 | 0 | 15 | 0.8484 | 4.7 |
| -40.6 | 0 | 18 | 0.5139 | 1.4 |
| -40.6 | 0 | 21 | 0.2868 | -1.7 |
| -40.6 | 3 | 21 | 0.1421 | -5.1 |
| -40.6 | 6 | 21 | 0.0745 | -8.1 |
| -40.6 | 9 | 21 | 0.0385 | -11.2 |
| -40.6 | 12 | 21 | 0.0201 | -14.2 |
| -40.6 | 15 | 21 | 0.0107 | -17.1 |
| -40.6 | 18 | 21 | 0.0056 | -20.4 |
| -40.6 | 21 | 21 | 0.0033 | -23.5 |
| -40.6 | 24 | 21 | 0.0022 | -26.4 |
| -40.6 | 27 | 21 | 0.0016 | -29.5 |
| -40.6 | 30 | 21 | 0.0014 | -32.7 |
| -40.6 | 30 | 24 | 0.0012 | -35.4 |
| -40.6 | 30 | 27 | 0.0012 | -37.87 |
| -40.6 | 30 | 30 | 0.0012 | -40.2 |

| Use Noise Source: Atlantic AS6333 |  |  |
| --- | --- | --- |
| Freq : 1.0 - 12.0GHz | Power Level: -41.8dBm |  |
| DC Supply: 28V |  |  |

| X |  |  |  |  |
| --- | --- | --- | --- | --- |
| Input Power [dBm]: | Attenuator Value A [dB]: | Attenuator Value: B [dB]: | Detector Value: | Power Meter Value [dBm]: |
| -41.8 | 0 | 0 | - | 11.5 |
| -41.8 | 0 | 3 | - | 10.6 |
| -41.8 | 0 | 6 | - | 9.2 |
| -41.8 | 0 | 9 | - | 7.2 |
| -41.8 | 0 | 12 | - | 4.7 |
| -41.8 | 0 | 15 | 0.8241 | 2.2 |
| -41.8 | 0 | 18 | 0.4526 | -1.0 |
| -41.8 | 0 | 21 | 0.2494 | -3.9 |
| -41.8 | 3 | 21 | 0.1289 | -6.9 |
| -41.8 | 6 | 21 | 0.0642 | -10.0 |
| -41.8 | 9 | 21 | 0.0332 | -12.9 |
| -41.8 | 12 | 21 | 0.0168 | -15.9 |
| -41.8 | 15 | 21 | 0.0095 | -18.6 |
| -41.8 | 18 | 21 | 0.0050 | -21.9 |
| -41.8 | 21 | 21 | 0.0030 | -25.0 |
| -41.8 | 24 | 21 | 0.0021 | -28.0 |
| -41.8 | 27 | 21 | 0.0016 | -30.7 |
| -41.8 | 30 | 21 | 0.0013 | -33.9 |
| -41.8 | 30 | 24 | 0.0013 | -36.5 |
| -41.8 | 30 | 27 | 0.0011 | -38.7 |
| -41.8 | 30 | 30 | 0.0011 | -40.8 |

### Detector Calibration and Attenuator Sweep Pol Y:

| Load configuration: Detector_Calibration_PAX.csa |  |  |
| --- | --- | --- |
| Freq : 4.0 GHz | Power Level: -20dBm | Sweep Time: -20sec |
| Sweep Mode: CW | Power Level at PAX Input (Measured): -40.6dBm |  |

| Y |  |  |  |  |
| --- | --- | --- | --- | --- |
| CW Input Power [dBm]: | Attenuator Value A [dB]: | Attenuator Value: B [dB]: | Detector Value: | Power Meter Value [dBm]: |
| -40.6 | 0 | 0 | - | 13.1 |
| -40.6 | 0 | 3 | - | 13.1 |
| -40.6 | 0 | 6 | - | 12.4 |
| -40.6 | 0 | 9 | - | 10.3 |
| -40.6 | 0 | 12 | 0.8471 | 7.4 |
| -40.6 | 0 | 15 | 0.7982 | 4.5 |
| -40.6 | 0 | 18 | 0.4339 | 1.2 |
| -40.6 | 0 | 21 | 0.2342 | -1.9 |
| -40.6 | 3 | 21 | 0.1141 | -5.3 |
| -40.6 | 6 | 21 | 0.0595 | -8.2 |
| -40.6 | 9 | 21 | 0.0305 | -11.3 |
| -40.6 | 12 | 21 | 0.016 | -14.3 |
| -40.6 | 15 | 21 | 0.0087 | -17.3 |
| -40.6 | 18 | 21 | 0.0046 | -20.6 |
| -40.6 | 21 | 21 | 0.0029 | -23.7 |
| -40.6 | 24 | 21 | 0.002 | -26.7 |
| -40.6 | 27 | 21 | 0.0015 | -29.8 |
| -40.6 | 30 | 21 | 0.0014 | -33.0 |
| -40.6 | 30 | 24 | 0.0013 | -35.7 |
| -40.6 | 30 | 27 | 0.0013 | -48.2 |
| -40.6 | 30 | 30 | 0.0012 | -40.5 |

| Use Noise Source: Atlantic AS6333 |  |  |
| --- | --- | --- |
| Freq : 1.0 - 12.0GHz | Power Level: -41.8dBm |  |
| DC Supply: 28V |  |  |

| Y |  |  |  |  |
| --- | --- | --- | --- | --- |
| Input Power [dBm]: | Attenuator Value A [dB]: | Attenuator Value: B [dB]: | Detector Value: | Power Meter Value [dBm]: |
| -41.8 | 0 | 0 | - | 11.8 |
| -41.8 | 0 | 3 | - | 11.0 |
| -41.8 | 0 | 6 | - | 9.6 |
| -41.8 | 0 | 9 | - | 7.8 |
| -41.8 | 0 | 12 | 0.8471 | 5.3 |
| -41.8 | 0 | 15 | 0.752 | 2.8 |
| -41.8 | 0 | 18 | 0.4066 | -0.3 |
| -41.8 | 0 | 21 | 0.2190 | -3.2 |
| -41.8 | 3 | 21 | 0.1111 | -6.2 |
| -41.8 | 6 | 21 | 0.0548 | -9.3 |
| -41.8 | 9 | 21 | 0.0281 | -12.2 |
| -41.8 | 12 | 21 | 0.0142 | -15.3 |
| -41.8 | 15 | 21 | 0.008 | -18.0 |
| -41.8 | 18 | 21 | 0.0043 | -21.4 |
| -41.8 | 21 | 21 | 0.0028 | -24.5 |
| -41.8 | 24 | 21 | 0.0020 | -27.6 |
| -41.8 | 27 | 21 | 0.0016 | -30.4 |
| -41.8 | 30 | 21 | 0.0014 | -33.8 |
| -41.8 | 30 | 24 | 0.0013 | -36.3 |
| -41.8 | 30 | 27 | 0.0012 | -38.5 |
| -41.8 | 30 | 30 | 0.0013 | -40.5 |

# Inspection:

### Visual:

| Fibre Connectors Clean | Comments: | OK NOK |
| --- | --- | --- |
| Fibre Connector Mechanical | Comments: | OK NOK |
| RF Cable | Comments: | OK NOK |
| RF Connectors Clean | Comments: | OK NOK |

### Functions:

| Read Temperature Value | Comments: | OK NOK |
| --- | --- | --- |
| LNA Settings | Comments: | OK NOK |
| Temperature Stabilisation | Comments: | OK NOK |
|  |  | OK NOK |